## Instructions
New Version January 2023
Do not change the format of the Acquisitions or Samples tab.Other tabs can be added and edited as needed.
When you save a sheet from bluesky, it will only contain the Bar and Acquisitions tab, with no instructions, so save this file somewhere for your reference
It is fine to just produce a sheet yourself with only these sheets and no data validation, and you are encouraged to do that if you are having problems
TEST your bar ahead of beamtime by using the "rsoxs_plans" package https://github.com/NSLS-II-SST/rsoxs_scans - which includes the most up to date spreadsheet sample and the same code which will be used at the beamline
accurate time estimates and breakdowns for each acquisition can be obtained using this package.
Instructions - READ CAREFULLY
visit https://wiki-nsls2.bnl.gov/beamline7ID1/index.php/RSoXS_Samples to see detailed instructions if needed
1.) Enter all sample info in the "Bar" tab - be as thurough as you are able and enter values for ALL bold columns. This information will follow the sample data, and will make analysis easier!
2.) Enter acquisitions in the "Acquisitions" tab, data validation is in place so you must chose a sample which you listed in the Bar tab and valid configurations and plans
Information about possible acquisitions and arguments are given on the wiki

## Bar
Parameter/ Index | bar_name | sample_id | sample_name | project_name | institution | proposal_id | bar_spot | front | grazing | angle | height | sample_desc | project_desc | sample_priority | notes | sample_set | components | composition | thickness | density | sample_date
Description | REQUIRED unique name for this bar | REQUIRED unique identifier for the sample. | REQUIRED will be in file name (readable by whoever is loading and measuring samples) | REQUIRED folder name for data The largest degree of sample classification | REQUIRED Short abbreviation of your institution that will be added to folder names. | REQUIRED 6 digit proposal for this measurement. This must match the approved proposal for this beamtime. General user proposals generally begin with GU- and contain 6 digits. | REQUIRED Where on the bar you loaded the sample rows 1-27 Column A-D. This will be used during alignment, so be careful! | REQUIRED Is the sample mounted on the flat side of the bar (front)? | REQUIRED Is the sample going to stop X-rays from penetrating? It is very important to indicate this as all calculations of angle will depend on this. | REQUIRED Incident angle where you want to collect data by default. This parameter is overwritten by acquisition plans that specify specific angles or angle ranges. | REQUIRED Distance from the surface of the sample bar to the top of the sample. | full text description of your sample. This will all be searchable in the database | Extra information about the project | 0 - critical 100-optional This can be used for organizing the acquisition order of the bar | Any information about the scans. Enter all Acquisitions in the Acquisitions tab, but you can describe in words what you are trying to measure here. Anything entered here will be searchable later in the database | optional sub-project grouping can help to sort data | optional list of components | optional quantitative componsition, such as chemical formula for NEXAFS analysis | optional Thickness of sample (nm) for absorption quantification | optional density of sample for absorption quantification | optional date when sample was created
Rules | all bar cells should share this exact name | Must match any physical sample labels | plain english is encouraged | any characters allowed on a linux filesystem | Be consistent across all beamtimes | Only the 6 digit number | be consistent see wiki for orientation help | TRUE OR FALSE | TRUE or FALSE | (transmission) 0 degrees = normal incidence, while 20 = 20 degrees from normal incidence, -20 - 20 degrees in the other direction from normal incidence. | Value in mm. |  |  | number between 0 and 100
Example | testbar | EG01 | P3HT-AN-120C | NEXAFS or RSoXS or OPVs | NIST or NCSU or UPENN | 310704 | 1A OR 12CD OR third sample from top, blue and shiny | TRUE | FALSE | (transmission) -180 (from back of sample) (grazing) 20 (20 degrees from grazing) | 0.25 mm | something about the sample | study of x effect on y | 1 | 0 | opvs | p3ht | C1234H342O2F65 | 0 | 0 | 0
Notes | 0 | this is what will be referenced in the acquisitions, so keep it simple | 0 | 0 | 0 | 0 | program will present you with this when you are identifying the samples in an image | 0 | If TRUE, any angle specified will be assumed to be a grazing angle (that is an angle with respect to the surface of the sample, where 90 degrees means normal to the sample). If FALSE, any angle specified will be assumed to be relative to the normal of the film (0 degrees is normal to the sample). | For transmission measurements (grazing = FALSE), the angle is taken with respect to the FRONT/BACK sample normal vector. The default is 180 degrees. For grazing measurements (grazing = True), specify the angle with respect to the front surface tangent. | Si Wafers are usually 250 microns thick (0.25 mm) | 0 | 0 | optional sorting method for queue creation | this is a nice place to explain in english what scans you want to run, which can later be translated into the acquisition tab | 0 | useful for seaching for similar samples taken previously | can be imported directly in QANT for Kramers Kronig | 0 | can be imported directly in QANT for Kramers Kronig | 0
1 | testbar | p_1 | p3ht | testing | NIST | 310704 | 1A | TRUE | FALSE | 180 | 0.25 | something about the sample | I'm doing science | 1 | I'd like to start by calling into question the whole concept of a polymer. What do we really know about these things?? | opvs | p3ht | C1234H342O2F65 | 125 | 1.2 | 2023-01-08
2 | testbar | p_2 | pcbm | testing | NIST | 310704 | 2B | TRUE | TRUE | 20 | 0.25 | soemthing special about this samples | I'm doing science | 2 | do a carbon scan at a bunch of different energies | opvs | pcbm | C60 | 10 | 1.5 | 2023-01-08

## Acquisitions
Parameter/ Index | sample_id | configuration | type | priority | edge | ratios | frames | repeats | speed | cycles | diameter | spiral_step | pol_mode | polarizations | angles | exposure_time | temperatures | temp ramp speed | wait at temperature | grating | diode_range | group | slack_message_start | slack_message_end
Description | Must exactly match a sample_id from the Bar sheet | Measurement Configuration | Type of measurement | Determines which order scans will be run, lowest value first. | Which elemental edge or energy ranges you wish to scan. | the ratio of resolution / speed between the different regions defined by the "edge" parameter | How many steps you want per energy scan (default 'full' which is 112 images, or specify the exact number you want) or how fast you want to run in ev/sec NEXAFS (default "normal" or 0.2), radius of the spiral (default 1.8) | How many images you want per step. | eV/sec to scan NEXAFS. | how many times to sweep energy up and back down | diameter in mm of spiral scan extent | step size for a spiral scan | when specifying a polarization for a NEXAFS scan, wether the specification is relative to the sample surface or relative to the lab | single angle or list of polarization angles to take measurement at | angles for rotation of the sample about the vertical axis | exposure time(s) in seconds for rsoxs scans and spirals | temperature(s) to move to between / during scans. | the temperature ramp speed deg C / min | wait for the temperature to reach the setpoint before starting the scans or not | which grating to use for measurements | high or low range for diode measurements | #N/A | #N/A | #N/A
Rules |  | Choose from WAXSNEXAFS, WAXS, SAXS, SAXSNEXAFS, SAXS_liquid, WAXS_liquid | Choose from RSoXS, NEXAFS, Spiral | Must be an integer from 1 to 100 | Must match an entry in the lookup table (see Notes), OR be a single energy OR be a list of energies within hard brackets | must match the number of regions defined by ratios. (length of ratios -1) for built in ratio tables, look up the length needed | this is only a estimate, the threshold energies will always be favored. See the examples in the jupyter notebook, and run dry runs there to see actual numbers of exposures and experimental times | Repeating exposures at a single step has much less overhead ~1 second than defining multiple steps at one energy ~4 seconds | built in speeds are good starting points. 0.1 is slow (5 minutes or so for a scan) 0.3 is fast (<2 minutes / scan) | any non negative integer |  |  | in lab frame 0 is horizontal in board, 90 vertical up in sample frame, 0 is normal to the sample surface (minimum possible is the grazing angle of the sample), 90 is in the plane of the sample | -1 -> circular -0.5 -> circular counter clockwise 0-180 -> angle from horizontal (sample frame: out of plane) through vertical (sample frame: in plane) | number or list of numbers to take scans | for rsoxs scans, a more complex nomenclature is available to define exposures for certain energy ranges. See the notebook example or the wiki | repeat the scans above for each temperature. | 0.1 - 100 are valid speeds | True/False | rsoxs (default), 1200, 250 | high, low | #N/A | #N/A | #N/A
Example | 0 | SAXS | Spiral | 2 | carbon OR 285 OR [270,280,290,400] | 5,1,5 -> go through the first and last region 5 times as fast (NEXAFS) or have the energy steps 5 times more spread out (RSoXS) than the central region | short | 5 -> repeat every exposure 5 times before moving to the next energy step' | fast | 0 -> no sweeping back down | 1.5 | 0.3 | 20 | 20,30,55,70,90 | 20,40,55,70,90 | 1 | 120 | 10 | True | rsoxs | high | #N/A | #N/A | #N/A
Notes | 0 | Determines which detector is used / slits | Depending on the value of this cell, unnecessary acquisition parameters will be greyed out and locked. | Normal way to sort queue steps | Current Lookup table entries are given here [LINK] | [11 empty] | ONLY used with the TEM holder - not valid for normal sample bar |  |  | WARNING - moving gratings between measurements might result in loss of energy calibration | high should be used in any case where saturation might be a possibility - the majority of measurements | #N/A | #N/A | #N/A
1 | p_1 | SAXS | rsoxs | 1 | carbon | [16 empty] | alignment | about to start alignment! | done with alignment
3 | p_2 | WAXS | spiral | 3 | 270 | [5 empty] | 2.1 | 0.3 | [9 empty] | 0 | starting with the data | tagging and messages from the spreadsheet are working!

## SheetRulesAndMetaData
Sheet Rules and Metadata | Users should not modify this sheet
Syntax
Sheet Version Number | 2023-1.1 | Cycle Active.Version
Sheet | Parameter | Description | Rules | Example | Notes
Bar | bar_name | REQUIRED unique name for this bar | all bar cells should share this exact name | testbar
Bar | sample_id | REQUIRED unique identifier for the sample. | Must match any physical sample labels | EG01 | this is what will be referenced in the acquisitions, so keep it simple
Bar | sample_name | REQUIRED will be in file name (readable by whoever is loading and measuring samples) | plain english is encouraged | P3HT-AN-120C
Bar | project_name | REQUIRED folder name for data The largest degree of sample classification | any characters allowed on a linux filesystem | NEXAFS or RSoXS or OPVs
Bar | institution | REQUIRED Short abbreviation of your institution that will be added to folder names. | Be consistent across all beamtimes | NIST or NCSU or UPENN
Bar | proposal_id | REQUIRED 6 digit proposal for this measurement. This must match the approved proposal for this beamtime. General user proposals generally begin with GU- and contain 6 digits. | Only the 6 digit number | 310704
Bar | bar_spot | REQUIRED Where on the bar you loaded the sample rows 1-27 Column A-D. This will be used during alignment, so be careful! | be consistent see wiki for orientation help | 1A OR 12CD OR third sample from top, blue and shiny | program will present you with this when you are identifying the samples in an image
Bar | front | REQUIRED Is the sample mounted on the flat side of the bar (front)? | TRUE OR FALSE | TRUE
Bar | grazing | REQUIRED Is the sample going to stop X-rays from penetrating? It is very important to indicate this as all calculations of angle will depend on this. | TRUE or FALSE | FALSE | If TRUE, any angle specified will be assumed to be a grazing angle (that is an angle with respect to the surface of the sample, where 90 degrees means normal to the sample). If FALSE, any angle specified will be assumed to be relative to the normal of the film (0 degrees is normal to the sample).
Bar | angle | REQUIRED Incident angle where you want to collect data by default. This parameter is overwritten by acquisition plans that specify specific angles or angle ranges. | (transmission) 0 degrees = normal incidence, while 20 = 20 degrees from normal incidence, -20 - 20 degrees in the other direction from normal incidence. | (transmission) -180 (from back of sample) (grazing) 20 (20 degrees from grazing) | For transmission measurements (grazing = FALSE), the angle is taken with respect to the FRONT/BACK sample normal vector. The default is 180 degrees. For grazing measurements (grazing = True), specify the angle with respect to the front surface tangent.
Bar | height | REQUIRED Distance from the surface of the sample bar to the top of the sample. | Value in mm. | 0.25 mm | Si Wafers are usually 250 microns thick (0.25 mm)
Bar | sample_desc | full text description of your sample. This will all be searchable in the database |  | something about the sample
Bar | project_desc | Extra information about the project |  | study of x effect on y
Bar | sample_priority | 0 - critical 100-optional This can be used for organizing the acquisition order of the bar | number between 0 and 100 | 1 | optional sorting method for queue creation
Bar | notes | Any information about the scans. Enter all Acquisitions in the Acquisitions tab, but you can describe in words what you are trying to measure here. Anything entered here will be searchable later in the database |  |  | this is a nice place to explain in english what scans you want to run, which can later be translated into the acquisition tab
Bar | sample_set | optional sub-project grouping can help to sort data |  | opvs
Bar | components | optional list of components |  | p3ht | useful for seaching for similar samples taken previously
Bar | composition | optional quantitative componsition, such as chemical formula for NEXAFS analysis |  | C1234H342O2F65 | can be imported directly in QANT for Kramers Kronig
Bar | thickness | optional Thickness of sample (nm) for absorption quantification
Bar | density | optional density of sample for absorption quantification |  |  | can be imported directly in QANT for Kramers Kronig
Bar | sample_date | optional date when sample was created
Acquisitions | sample_id | Must exactly match a sample_id from the Bar sheet
Acquisitions | configuration | Measurement Configuration | Choose from WAXSNEXAFS, WAXS, SAXS, SAXSNEXAFS, SAXS_liquid, WAXS_liquid | SAXS | Determines which detector is used / slits
Acquisitions | type | Type of measurement | Choose from RSoXS, NEXAFS, Spiral | Spiral | Depending on the value of this cell, unnecessary acquisition parameters will be greyed out and locked.
Acquisitions | priority | Determines which order scans will be run, lowest value first. | Must be an integer from 1 to 100 | 2 | Normal way to sort queue steps
Acquisitions | edge | Which elemental edge or energy ranges you wish to scan. | Must match an entry in the lookup table (see Notes), OR be a single energy OR be a list of energies within hard brackets | carbon OR 285 OR [270,280,290,400] | Current Lookup table entries are given here [LINK]
Acquisitions | ratios | the ratio of resolution / speed between the different regions defined by the "edge" parameter | must match the number of regions defined by ratios. (length of ratios -1) for built in ratio tables, look up the length needed | 5,1,5 -> go through the first and last region 5 times as fast (NEXAFS) or have the energy steps 5 times more spread out (RSoXS) than the central region
Acquisitions | frames | How many steps you want per energy scan (default 'full' which is 112 images, or specify the exact number you want) or how fast you want to run in ev/sec NEXAFS (default "normal" or 0.2), radius of the spiral (default 1.8) | this is only a estimate, the threshold energies will always be favored. See the examples in the jupyter notebook, and run dry runs there to see actual numbers of exposures and experimental times | short
Acquisitions | repeats | How many images you want per step. | Repeating exposures at a single step has much less overhead ~1 second than defining multiple steps at one energy ~4 seconds | 5 -> repeat every exposure 5 times before moving to the next energy step'
Acquisitions | speed | eV/sec to scan NEXAFS. | built in speeds are good starting points. 0.1 is slow (5 minutes or so for a scan) 0.3 is fast (<2 minutes / scan) | fast
Acquisitions | cycles | how many times to sweep energy up and back down | any non negative integer | 0 -> no sweeping back down
Acquisitions | diameter | diameter in mm of spiral scan extent |  | 1.5
Acquisitions | spiral_step | step size for a spiral scan |  | 0.3
Acquisitions | pol_mode | when specifying a polarization for a NEXAFS scan, wether the specification is relative to the sample surface or relative to the lab | in lab frame 0 is horizontal in board, 90 vertical up in sample frame, 0 is normal to the sample surface (minimum possible is the grazing angle of the sample), 90 is in the plane of the sample | 20
Acquisitions | polarizations | single angle or list of polarization angles to take measurement at | -1 -> circular -0.5 -> circular counter clockwise 0-180 -> angle from horizontal (sample frame: out of plane) through vertical (sample frame: in plane) | 20,30,55,70,90
Acquisitions | angles | angles for rotation of the sample about the vertical axis | number or list of numbers to take scans | 20,40,55,70,90
Acquisitions | exposure_time | exposure time(s) in seconds for rsoxs scans and spirals | for rsoxs scans, a more complex nomenclature is available to define exposures for certain energy ranges. See the notebook example or the wiki | 1
Acquisitions | temperatures | temperature(s) to move to between / during scans. | repeat the scans above for each temperature. | 120 | ONLY used with the TEM holder - not valid for normal sample bar
Acquisitions | temp ramp speed | the temperature ramp speed deg C / min | 0.1 - 100 are valid speeds | 10
Acquisitions | wait at temperature | wait for the temperature to reach the setpoint before starting the scans or not | True/False | True
Acquisitions | grating | which grating to use for measurements | rsoxs (default), 1200, 250 | rsoxs | WARNING - moving gratings between measurements might result in loss of energy calibration
Acquisitions | diode_range | high or low range for diode measurements | high, low | high | high should be used in any case where saturation might be a possibility - the majority of measurements